Deviation Report

# Report Information

|  |  |  |  |
| --- | --- | --- | --- |
| Title | Lasair particle monitoring equipment stopped recording results after first 400L sample | Priority | Major |
| Status | In Progress | Type | Equipment |
| Date | 03/21/2025 | Department | QA |
| Batch/Lot | 10000062 | Quantity | entire batch |
| Planned | No | Initiator | Matt |
| Record ID | DEV\_20250808\_032026 | Generated | 08/08/2025 03:21 AM |

# 👤 User Input - Original Deviation Description

On 21MAR2025, the Lasair particle monitoring equipment and stopped recording results after the first 400L sample was taken on the data printout. Review of the instrument shows that it sampled continuously from 0838 to 1631 on 21MAR2025, however the print out states "invalid" with no alerts or alarms with no particle quantification.

# 🤖 AI-Generated Analysis